Tender Specifications

- FACO. Power Supply:

Fixed outputs DC: + 5 V, ± 12 V, 1 A.

Variable outputs DC: ± 12 V, 0.5 A.

AC output: 12 V or 24 V.

Connection to the PLC-BS modules through 25 pin Sub-D connectors.

LED’s voltage indicators.

Robust construction.

- BS5. Ovens Test Module:

Painted steel box. Connection diagrams for each transducer are represented graphically.

Oven chamber.

Heating element:

Oven heating element made up of two parallel resistances with a maximum dissipation power of 500 W. The heating element power supply is of 0 – 30 V AC. Inside the heating element there is a temperature sensor element.

Fan:

Fan with changeable speed that can be operated varying the fan energy supply voltage.

Energy supply voltage: +12 Vdc (max). Maximum power: 0.96 W. Maximum air flow: 2.5 l/s.

Thermocouples:

Four thermocouples placed inside the oven, each one of them at a different height. Temperature range: -184 ºC – 400 ºC.

Platinum resistance thermometer:

Platinum resistance temperature detector, suitable for measuring air and gas temperatures. Temperature range: -70 ºC – 600 ºC. Resistance (0 ºC): 100+/-0.1 Ω.

Thermistor:

NTC thermistor for temperature measurement and control, with great sensitivity and stability. Resistance at 25 ºC: 5.8 KΩ. Temperature range: -40 ºC – 125 ºC.

Semiconductor temperature sensor:

Reverse polarized diode. The current through the diode depends on the temperature at which balance with the surrounding environment is achieved. Therefore it needs a conditioning circuit able to transform this current variation in voltage proportional to temperature.

Feedback sensor for PLC control:

Signal conditioning circuit of a temperature sensor.

Voltage output range: 0 – 10 Vdc.

Cables and Accessories, for normal operation.

Manuals: This unit is supplied with the following manuals: Required Services, Assembly and Installation, Starting-up, Safety, Maintenance & Practices Manual.

Each module may operate independently of one another.

Dimensions: 405 x 300 x 470 mm approx. (15.94 x 11.81 x 18.50 inches approx.) Weight: 7.5 Kg approx. (16.53 pounds approx.)

EXERCISES AND PRACTICAL POSSIBILITIES

1.-  Heating element. Raise the oven internal temperature, over the environmental temperature, using a heating element to do tests and practices related with temperature measurement.

2.-  To use a fan as refrigerating element of the oven.

3.-  To use thermocouples as temperature sensors elements inside the oven. Temperature measurement using a thermocouple.

4.-  To measure temperature inside the oven using a platinum resistance thermometer.

5.-  To measure temperature inside the oven using a thermistor temperature sensor.

6.-  Temperature measurement using a thermistor, based on its negative temperature coefficient.

7.-  To obtain the temperature value inside the oven using a semiconductor sensor (diode).

8.-  Using the analog inputs and outputs of the PLC to read and write analog signals.

9.-  Read the temperature sensor signal with the PLC.

10.-Manual control of the actuator through the PLC.

11.-PID control of the temperature value through the PLC.

12.-Observe the effects of the temperature control to changes in the PID parameters.

Optional

PLC-BS5/ICAI. Interactive Computer Aided Instruction Software System:

This complete software package consists of an Instructor Software (EDIBON Classroom Manager - ECM-SOF) totally integrated with the Student Software (EDIBON Student Labsoft - ESL-SOF). Both are interconnected so that the teacher knows at any moment what is the theoretical and practical knowledge of the students.

- ECM-SOF. EDIBON Classroom Manager (Instructor Software).

ECM-SOF is the application that allows the Instructor to register students, manage and assign tasks for workgroups, create own content to carry out Practical Exercises, choose one of the evaluation methods to check the Student knowledge and monitor the progression related to the planned tasks for individual students, workgroups, units, etc... so the teacher can know in real time the level of understanding of any student in the classroom.

Innovative features:

• User Data Base Management.

• Administration and assignment of Workgroup, Task and Training sessions.

• Creation and Integration of Practical Exercises and Multimedia Resources.

• Custom Design of Evaluation Methods.

• Creation and assignment of Formulas & Equations.

• Equation System Solver Engine.

• Updatable Contents.

• Report generation, User Progression Monitoring and Statistics.

- ESL-SOF. EDIBON Student Labsoft (Student Software).

ESL-SOF is the application addressed to the Students that helps them to understand theoretical concepts by means of practical exercises and to prove their knowledge and progression by performing tests and calculations in addition to Multimedia Resources. Default planned tasks and an Open workgroup are provided by EDIBON to allow the students start working from the first session. Reports and statistics are available to know their progression at any time, as well as explanations for every exercise to reinforce the theoretically acquired  technical knowledge.

Innovative features:

• Student Log-In & Self-Registration.

• Existing Tasks checking & Monitoring.

• Default contents & scheduled tasks available to be used from the first session.

• Practical Exercises accomplishment by following the Manual provided by EDIBON.

• Evaluation Methods to prove your knowledge and progression.

• Test self-correction.

• Calculations computing and plotting.

• Equation System Solver Engine.

• User Monitoring Learning & Printable Reports.

• Multimedia-Supported auxiliary resources.